FLC_PHY2 Front end chip

DATASHEET

# Table of content

Table of content	2

Table of figures	3

Figures	3

Tables	3

Equation	4

General description	5

One channel description	7

Charge preamplifier	7

Shapers	11

Track and hold	11

Output buffer	13

Multiplexer description	13

Pin definition	15

Characteristics & measurement	17

DC Levels	17

Crosstalk	18

Noise	19

Transient	22

Linearity	24

Linearity measurement – Cf=3pF	25

Linearity measurement – Cf=1.6pF	26

Linearity measurement – Cf=0.8pF	27

Linearity measurement – Cf=0.4pF	28

Linearity measurement – Cf=0.2pF	29

Pedestal dispersion	30

ANNEX 1 – FLC_PHY2 pinout, bias and current	31

ANNEX 2 – FLC_PHY2 versus FLC_PHY1 pinout	32

ANNEX 3 - Preamp schematic	33

ANNEX 4 – shaper gain 1 schematic	34

ANNEX 5 – Shaper gain 10 schematic	35

ANNEX 6 – Shapers OPAMP schematic	36

ANNEX 7 – Widlar schematic	37

# Table of figures

## Figures

Figure 1 General block schema of FLC_PHY2	5

Figure 2 - Block schema of one channel	7

Figure 3 - Schema of charge preamplifier	8

Figure 4 - first stage of charge preamplifier	9

Figure 5 - preamplifier resistor feedback	10

Figure 6 - CRRC shaper schematic	11

Figure 7- Block diagram of the track and hold	12

Figure 8 - Block diagram of the shift register	14

Figure 9 - Chronogram of the multiplexed readout	14

Figure 10 - ENC=f(Cd) with Cf=1.6pF	19

Figure 11 - Output noise vs preamp gain	20

Figure 12 - output noise versus feedback capacitance	21

Figure 13 - Output noise vs Input PMOS current bias	21

Figure 14 - Output gain 1 versus feedback capacitance	22

Figure 15 - Measured feedback capa versus theoretic capa	23

Figure 16 - Widlar & Common collector transient for comparison	23

Figure 17 - Linearity - Cf=3pF	25

Figure 18 - Linearity - Cf=1.6pF	26

Figure 19 - Linearity - Cf=0.8pF	27

Figure 20 - Linearity - Cf=0.4pF	28

Figure 21 - Linearity - Cf=0.2pF	29

Figure 22 - Pedestal measurement on Gain 1 output	30

Figure 23 - Pedestal measurement on gain 10 output	30

## Tables

Table 1- Compensation capacitance values	9

Table 2 - Gain versus feedback capacitance switch configuration	10

Table 3 - follower switch SW description	12

Table 4 - measured vs simulated DC levels	17

Table 5 - Crosstalk measurement	18

Table 6 - ENC vs Cd, Cf=1.6pF	19

Table 7 - Noise versus shut down stages	20

Table 8 - Measured gains and peaking times	22

Table 9 - Characterization Cf=3pF	25

Table 10 - Characterization Cf=1.6pF	26

Table 11 - Characterization Cf=0.8pF	27

Table 12 - Characterization Cf=0.4pF	28

Table 13 - Characterization Cf=0.2pF	29

Table 14 - Pedestal measurement results	30

## Equation

Equation 1 - Shaper transfer function	11

# General description

The FLC_PHY2 Front-end chip is an 18-channel charge input front end circuit. It provides a shaped signal proportional to the input charge.

Buffer

OUT G=1

Channel 0

Shaper

G=1

T&Hold

IN 0

Shaper

G=10

T&Hold

4

Gain select

Channel 1

Shaper

G=1

T&Hold

IN 1

Shaper

G=10

T&Hold

Channel 17

preamp

Shaper

G=1

T&Hold

IN 17

Shaper

G=10

T&Hold

Buffer

OUT G=10

preamp

preamp

Figure 1 General block schema of FLC_PHY2

Each channel is made of a variable-gain charge preamplifier followed by two parallel shaping filter of different gain (1 and 10) using a CRRC structure. Each of these shaper are followed by a track & hold device driving two single multiplexed output (one for the unit gain and one for the gain 10).

The bias of each stage is common for every channel.

Figure 1 is a simplified schema block of the whole chip.

# One channel description

Amp

OPA

OPA

MUX out Gain=1

MUX out Gain=10

Input

Figure 2 - Block schema of one channel

Figure 2 shows a simplified block schema of one channel. The first stage is made of a multi-gain charge preamplifier. Four switchable feedback capacitance allow to tune the gain within 16 gains. After the preamplifier, the circuit is doubled to optimise dynamic range. Two shapers with different gains (1 and 10) allows to reduce noise. Hence there is two shaped signal available. Choosing one or the other may depend on the magnitude of the signal. A track and hold device allowing multiplexing follows each of the shaper.

Each functional block will be described in detail below.

## Charge preamplifier

25

25

1

1

0.2pF

0.4pF

0.8pF

1.6pF

1pF

2pF

4pF

4pF

Figure 3 - Schema of charge preamplifier

The first stage of that preamplifier is a common source transistor biased by a current source made of two NMOS transistor connected as a current mirror.

It is recommended to flow at least 500µA in the input PMOS to reduce series noise by increasing the PMOS gm. That means Vbiasi_pa should be connected to the Vdd with a resistor smaller than 10kΩ (typ. 6.7kΩ for Ids=500µA or 3.3 kΩ for Ids=1mA).A compromise have to be done at this point by the designer between consumption and noise. Vbiasi_pa is the current source of the input PMOS and the bipolar cascode. The current flowing in the input PMOS can be calculated by subtracting the cascode current (provided by the Vbiasm_pa current source) to the input current source.

The noise can be reduced a little bit more by increasing the PMOS bulk voltage (pin Vdda). It is recommended to bias that bulk 2V higher than the high supply rail (Figure 4).

IN

Vbiasi_pa

Vdda

Figure 4 - first stage of charge preamplifier

The second stage is a common base bipolar transistor (NPN) in a cascode structure. The Vb pin allows to tune the input PMOS Vds. The Vbiasm_pa bias the bipolar transistor with a DC current.

After the second stage are the compensation capacitors that  allows to slow the preamplifier down in case of oscillation, if the impedance on the input is not enough capacitive. The value of the compensation capacitance are given in Table 1.

| Capacitance | Value |
| --- | --- |
| Switch 1 | 4pF |
| Switch 2 | 4pF |
| Switch 3 | 2pF |
| Switch 4 | 1pF |

Table 1- Compensation capacitance values

The third stage is a follower (common drain PMOS) biased by a current source (Vbiaso_pa).

The feedback capacitor is tuneable. 4 switched capacitors allows to choose within 16 gain. The different gain are listed in Table 2.

| Switch configuration |  |  |  | Feedback Capacitance | Gain |
| --- | --- | --- | --- | --- | --- |
| Capa SW0=1.6pF | Capa SW1=0.8pF | Capa SW2=0.4pF | Capa SW3=0.2pF |  |  |
| OFF | OFF | OFF | OFF | 0pF | N/A |
| OFF | OFF | OFF | ON | 0.2pF | 5V/pC |
| OFF | OFF | ON | OFF | 0.4pF | 2.5V/pC |
| OFF | OFF | ON | ON | 0.6pF | 1.67V/pC |
| OFF | ON | OFF | OFF | 0.8pF | 1.25V/pC |
| OFF | ON | OFF | ON | 1pF | 1V/pC |
| OFF | ON | ON | OFF | 1.2pF | 830mV/pC |
| OFF | ON | ON | ON | 1.4pF | 714mV/pC |
| ON | OFF | OFF | OFF | 1.6pF | 625mV/pC |
| ON | OFF | OFF | ON | 1.8pF | 555mV/pC |
| ON | OFF | ON | OFF | 2pF | 500mV/pC |
| ON | OFF | ON | ON | 2.2pF | 455mV/pC |
| ON | ON | OFF | OFF | 2.4pF | 416mV/pC |
| ON | ON | OFF | ON | 2.6pF | 385mV/pC |
| ON | ON | ON | OFF | 2.8pF | 357mV/pC |
| ON | ON | ON | ON | 3pF | 333mV/pC |

Table 2 - Gain versus feedback capacitance switch configuration

The resistor feedback is made of two unbalanced current mirror and a resistor (Figure 3).That makes the resistor virtually bigger.

The output transistor of the mirror is 25 times bigger than the input one, that means the current in the input line is 25 times lower. Considering there is two stage, the input current is 625 smaller than the output one which is flowing in the feedback resistor. The resistor seen at the input of the preamp is then virtually 625 bigger than its real value (36kΩ*625=22.5MΩ, Figure 5).

Preamp out

Preamp in

I

I/25

I/625

25

25

1

1

36kΩ

Figure 5 - preamplifier resistor feedback

## Shapers

Shapers are  CRRC active filters that use operational amplifiers (Figure 6). The differential input of an operational amplifier allows to reduce the pedestal dispersion of shapers due to a DC loop feedback.

The DC reference allows to tune the output DC level. The transfer function of shapers is given on Equation 1.

C1

R1

C2

R2

-

+

IN

DC ref

OUT

Figure 6 - CRRC shaper schematic

H(p)=−R2R1⋅R1C1p1+R1C1p1+R2C2p

Equation 1 - Shaper transfer function

## Track and hold

The track and hold allows to memorize an analogue value in a capacitance using a CMOS switch driven by the hold signal (pin H). That hold signal have to be synchronized with the peaking time of the shaper to ensure a maximum dynamic range and offset swing. The analogue value can then be red through the read CMOS switch. A follower forbids the charge of the capacitance to go away through the readout electronic. The value is then conserved during the reading. It is possible to choose between two different follower architecture. The first one is a simple common collector as on FLC_PHY1 chip. The second one use a differential structure (Widlar) to optimise pedestal dispersion. The pin SW permits to choose between common collector or Widlar structure (Table 3).

| SW=1(Vdd) | SW=0(Vss) |
| --- | --- |
| Common collector | Widlar structure |

Table 3 - follower switch SW description

The read switch is driven by a D latch cascaded as a shift register. That system allows to read sequentially the analogue value of the 2*18 channels with the multiplexed output. The first 18 read values are respectively the 18 gain-1 analogue value, while these channels are read, the gain-10 output is in a high impedance state. The 18 next read values are the 18 gain-10 analogue values while these channels are read, the gain-1 output is in a high impedance state. A block diagram of the track and hold is shown on Figure 7.

Reset

CMOS switch

CMOS switch

To Multiplexed output buffer

D latch

Clock

Reset

D

Q

From shaper

Hold

Clock

To D of channel N+1

From Q of channel N-1

Widlar

follower

IN

OUT

Follower selec. switch

Figure 7- Block diagram of the track and hold

## Output buffer

The output buffer is made of an OTA with a unit feedback on the negative input. The bias_buf pin is to bias the input differential peer of the OTA. 100 µA DC current is needed to bias that first stage. A 47 kΩ resistor to Vdd allows to obtain that current.

## Multiplexer description

The multiplexer allows to serialize the 2*18 outputs. This is made by enabling the read of the track and hold (see Figure 7) of each channel sequentially one after another. The read of each channel is controlled by the output of a D-latch. To read sequentially outputs, it is needed to provide a “1” at the input of the shift register (pin R), which is the D of the channel 0 of Gain 1 output. After a rising clock, the channel 0 of gain 1 output is enable (i.e. readable on the output). The shift register input have then to be set to “0” before the next rising clock to avoid having several “1” propagating in the shift register, therefore several channel available at the output at the same moment. After a rising clock, the channel 0 will be disabled and the channel 1 will be enable as the output of the channel 1 D-latch will be “1”. The shift register will then shift the enabled output up to the 17th channel of the gain 1 output. The next rising clock will then enable channel 0 of gain 10 output, and so on.When the 17th channel of gain 10 output will be enabled, the output register (pin Q_R) will be up. That output allows to control that the shift register propagate properly the enabling value and permits to cascade several chip. The block diagram of the shift register is on Figure 8.The chronogram is available on Figure 9.

Depending on measurement, next version may include parallel output sequence for gain 1 and gain 10 to ensure full logical compatibility with FLC_PHY1.

CLK

D

RST

SRIN

G1

READ0

G1

READ1

G1

READ2

G1

READ16

G1

READ17

D

D

D

D

CLK

D

RST

SROUT

D

D

D

D

G10

READ0

G10

READ1

G10

READ2

G10

READ16

G10

READ17

Figure 8 - Block diagram of the shift register

RST

CLK

SRIN

OUT 1

SROUT

G1

Ch.0

G1

Ch.1

G1

Ch.2

G1

Ch.3

Hold

0

1

2

3

34

35

36

OUT 10

G10

Ch.16

G10

Ch.17

Figure 9 - Chronogram of the multiplexed readout

# Pin definition

Ibm_pa

Rst_R

Ibo_sh10

SW

Ibi_wid

Ib_cc

Ibo_wid

Vddd

OUT_g10

Vdd_ota+diode

SW_C2

SW_C3

Vss_ota

SW_C0

SW3

Ib_ota

SW2

SW_C1

Vdda

In 0

Ibi_pa

Vb_pa + Vbf

Vss_pa

If_pa

Ibo_pa

Ibo_sh1

Ibi_sh1

H

D_R

Vss_setH

Ibi_sh10

In 16

Ck_R

In 1

In 2

In 3

In 4

In 5

In 6

In 7

In 8

Vss

In 9

In 10

In 11

In 12

In 15

In 14

In 13

SW0

In 17

Vdd_pa

SW1

Vss_pa

Vss_sh1

Vref_sh1

Q_R

Vdd_sh1

OUT_g1

Vss_sh10

Vref1_sh10

Vdd_sh10

Vref2_sh10

Vdd_setH

1

2

3

4

5

6

7

8

10

9

11

12

13

14

15

16

17

18

19

20

21

24

22

23

31

30

29

28

27

26

25

40

39

38

32

37

36

35

34

33

47

46

45

44

43

42

41

51

50

49

48

56

57

58

59

60

61

62

63

64

55

54

53

52

FLC_PHY2

TQFP 64 package

| Pin number | Pin name | definition |
| --- | --- | --- |
| 64, 1 – 8, 10 – 18 | In 0 – in 17 | 18 charge inputs |
| 9, 22, 23, 28, 41, 51, 60 | Vss | Low supply (substrate bias) |
| 19, 26, 31, 32, 40, 42 | Vdd | High supply |
| 63 | Vdda | Input transistor bulk bias |
| 20,21, 34, 35 | SW1, SW2, SW3, SW4 | CMOS preamp feedback capacitor switch SW0 : 1.6 pF SW1 : 0.8 pF SW2 : 0.4 pF SW3 : 0.2 pF |
| 36 – 39 | SW_C0, SW_C1, SW_C2, SW_C3 | CMOS preamp compensation capacitor switch SW_C0 : 4pF SW_C1 : 4pF SW_C2 : 2pF SW_C3 : 1pF |
| 47 | SW | Widlar or common collector buffer switch |
| 50 | Rst_R | CMOS shift register reset |
| 53 | Ck_R | CMOS shift register clock |
| 54 | H | CMOS Hold command |
| 52 | D_R | CMOS shift register input |
| 25 | Q_R | CMOS shift register output |
| 27 | OUT_g1 | Gain 1 multiplexed output |
| 43 | OUT_g10 | Gain 10 multiplexed output |
| 24 | Vref_sh1 | Gain 1 shaper voltage reference for differential |
| 29 | Vref1_sh10 | Gain 10 shaper voltage reference for differential |
| 30 | Vref2_sh10 | Gain 10 shaper voltage reference for anti-saturation |
| 62 | Ibi_pa | Preamplifier input stage current bias |
| 58 | Ibm_pa | Preamplifier middle-stage current bias |
| 57 | Ibm_pa | Preamplifier out-stage current bias |
| 61 | Vb_pa+Vbf | Base cascode transistor voltage bias and feedback resistor cascode transistor voltage bias |
| 59 | If_pa | Feedback preamplifier leakage current bias (should be around 100nA) |
| 56 | Ibi_sh1 | Current bias of gain 1 shaper input stage |
| 55 | Ibo_sh1 | Current bias of gain 1 shaper output stage |
| 49 | Ibi_sh10 | Current bias of gain 10 shaper input stage |
| 48 | Ibo_sh10 | Current bias of gain 10 shaper output stage |
| 46 | Ibi_wid | Current bias of widlar buffer input stage |
| 44 | Ibo_wid | Current bias of widlar buffer output stage |
| 45 | Ib_cc | Current bias of common collector buffer |
| 33 | Ib_OTA | Current bias of OTA output buffer |

# Characteristics & measurement

## DC Levels

Following DC levels have been measured and simulated on FLC_PHY2 chip :

| Block | Bias name | Simulation result | Measurement | Current yield |
| --- | --- | --- | --- | --- |
| Preamplifier | Ibi_pa | -3.3V | -3.3V | 1mA |
|  | Ibm_pa | -1.37V | -1.6V | 70uA |
|  | Ibo_pa | -1.28V | -1.6V | 100uA |
|  | If_pa | -0.7V | -0.9V | 100nA |
|  | Vb_pa | -3.8V | -3.8V | None – voltage ref. |
| Shaper 1 | Ibi_sh1 | -3.7V | -3.8V | 50uA |
|  | Ibo_sh1 | -3.7V | -3.8V | 50uA |
|  | Vref_sh1 | -3.8V | -3.8V | None – voltage ref. |
| Shaper 10 | Ibi_sh10 | -3.7V | -3.8V | 50uA |
|  | Ibo_sh10 | -3.7V | -3.8V | 50uA |
|  | Vref1_sh10 | -3.8V | -3.8V | None – voltage ref. |
|  | Vref2_sh10 | -2V | -2V | None – voltage ref. |
| Track&hold | Ib_cc | -1.4V | -1.5V | 50uA |
|  | Ibi_wid | -3.7V | -3.8V | 50uA |
|  | Ibo_wid | -3.7V | -3.7V | 50uA |
| OTA | Ib_ota | -3.7V | -4V | 50uA |

Table 4 - measured vs simulated DC levels

## Crosstalk

| Channel 7 |  | Signal magnitude @ peaking time : 1.364mV (1.48‰ of signal) |
| --- | --- | --- |
| Channel 8 |  | Signal magnitude @ peaking time : 1.568mV (1.69‰ of signal) |
| Channel 9 |  | HIT CHANNEL Injected charge : 5pC Signal magnitude @ peaking time : 924.7mV (100% of signal) |
| Channel 10 |  | Signal magnitude @ peaking time : V (0. 762‰ of signal) |
| Channel 11 |  | Signal magnitude @ peaking time : 639μV (0.691‰ of signal) |

Table 5 - Crosstalk measurement

The measured crosstalk is well beyond application expectation as shown on Table 5. Average crosstalk is below 1‰ in the 4 nearest channel around hit channel.

## Noise

As FLC_PHY2 does not include testpoint after preamplifier, it is more difficult to analyze noise. The only signal that can be monitored is the multiplexed output. It is thus not possible to sweep peaking time shaping to fit ENC measurement and extract serial noise.

Noise measurement made so far are below expectations. In order to compare noise with FLC_PHY1, the 1.6pF feedback capacitor will be selected for noise measurement. A 500uA input bias is set up.

ENC is calculated for different detector capacitors Cd. Results are available in Table 6.

| Cd (pF) | ENC (e-) |
| --- | --- |
| 0 | 3273 |
| 10 | 3496 |
| 18 | 3725 |
| 33 | 4250 |
| 68 | 5188 |

Table 6 - ENC vs Cd, Cf=1.6pF

Figure 10 - ENC=f(Cd) with Cf=1.6pF

By fitting ENC=f(Cd) serial noise is estimated around 2.42 nV/sqrt(Hz) with a parasitic capacitance of 112.77 pF. That estimation is physically impossible thus obtain such a big parasitic capacitance is not possible (parasitic capacitance should be around 15 pF). That means there is a wrong approximation in the calculation.

The end of that chapter will deal with measurement made to figure out why the noise is so high.

Theory assume that preamplifier is the major noise source, next stages are supposed to be not noisy. First work will be to slice off different stage to isolate noise sources.

First measurement shows up that the noise sounds to come from shaper rather than preamplifier. All stage of a channel have been shut down one after another. RMS noise have then be measured at the output of the whole channel. Results are presented in Table 7.

| Stage down | Measurement of | Output RMS noise |
| --- | --- | --- |
| OTA | Board buffer | 67.7 μV |
| Common collector | OTA+board buffer | 72.12 μV |
| Shaper 1 | OTA+board buf. + T&H | 77 μV |
| Preamp | OTA+board buf. + T&H + shaper | 184.88 μV |
|  | Whole line | 187.93 μV |

Table 7 - Noise versus shut down stages

It sounds like major noise source comes from the shaper stage. Some further measurements have been made with gain 10 shaper. Results are around twice better.

These bad results may come from very high resistor (33k and 100k) in the shaper. It will be necessary to analyze more in detail these noise results to improve results of next iteration.

Some parameter of the preamp have then be swept to try to minimize noise and improve knowledge of that noise. A commenter avec Christophe

Figure 11 - Output noise vs preamp gain

Figure 12 - output noise versus feedback capacitance

Figure 13 - Output noise vs Input PMOS current bias

## Transient

Figure 14 - Output gain 1 versus feedback capacitance

Shape of the output signal, gain and peaking time will be studied In this chapter. The aim is to determine that peaking time is constant when changing gain or buffer, etc. Peaking time versus magnitude will not be figured out as it is implicit in the linearity study.

As shown on Figure 14, the peaking time seems not to change when the gain is changing. Table 8 gives more details about peaking time for every gain. The dispersion of peaking time is small enough according to the shape of the output signal (Figure 14 ) to assume that it is not changing versus preamplifier gain.

| Theoretic Feedback capa | Theoretic gain | Measured gain | Peaking time Meas. |
| --- | --- | --- | --- |
| 0.2pF | 5V/pC | 5.61 V/pC | 217.84 ns |
| 0.4pF | 2.5V/pC | 2.85 V/pC | 217.37 ns |
| 0.6pF | 1.67V/pC | 1.92 V/pC | 215.39 ns |
| 0.8pF | 1.25V/pC | 1.45 V/pC | 214.7 ns |
| 1pF | 1V/pC | 1.17 V/pC | 214.7 ns |
| 1.2pF | 830mV/pC | 973 mV/pC | 212.92 ns |
| 1.4pF | 714mV/pC | 832 mV/pC | 212.46 ns |
| 1.6pF | 625mV/pC | 735 mV/pC | 212.77 ns |
| 1.8pF | 555mV/pC | 656 mV/pC | 213.47 ns |
| 2pF | 500mV/pC | 593 mV/pC | 214.21 ns |
| 2.2pF | 455mV/pC | 543 mV/pC | 216.44 ns |
| 2.4pF | 416mV/pC | 490 mV/pC | 215.45 ns |
| 2.6pF | 385mV/pC | 457 mV/pC | 216.61 ns |
| 2.8pF | 357mV/pC | 424 mV/pC | 216.76 ns |
| 3pF | 333mV/pC | 395 mV/pC | 215.45 ns |

Table 8 - Measured gains and peaking times

As gain is conditioned by feedback capacitance and assuming that MMIC capacitance has around 20% precision, we shall assume that gain will have a 20% precision. Measurement made on gain is shown on Figure 15. Measured capacitance are smaller than expected capacitance.

Figure 15 - Measured feedback capa versus theoretic capa

To ensure full compatibility with read-out part wathever the final choice on the buffer will be, peaking time has to be constant between Widlar structure and common collector buffer. Figure 16 shows that shape are exactly the same. There is a little gain modification that will be quantified in the linearity part but this is not critical for compatibility.

Figure 16 - Widlar & Common collector transient for comparison

## Linearity

Linearity has been measured for 5 different gains. An outlook of performance is available in the 5 characterisation sheet below. Linearity have been measured for both common collector and Widlar structure track and hold buffer. Results show up that Widlar structure is better in term of voltage swing. Linearity is only fraction of percent whatever the gain or the buffer is. Linearity results are beyond expectations. All curves have been taken with a 50 Ohm load. That means voltages are divided by 2 in every figure or table.

### Linearity measurement – Cf=3pF

#### Switch configuration

| Switch 0 | Switch 1 | Switch 2 | Switch 3 |
| --- | --- | --- | --- |
| ON | ON | ON | ON |

#### Measurement results

Figure 17 - Linearity - Cf=3pF

| Buffer Structure | Widlar | Common collector |
| --- | --- | --- |
| Qin max | 6.34 pC | 4.83 pC |
| Vout max | 1.26 V (over 50 ) | 0.94 V (over 50 ) |
| Gain | 199.2 mV/pC | 193 mV/pC |
| Feedback cap | 2.51 pF | 2.59pF |
| Non-linearity | 3.9‰ | 2.1‰ |

Table 9 - Characterization Cf=3pF

### Linearity measurement – Cf=1.6pF

#### Switch configuration

| Switch 0 | Switch 1 | Switch 2 | Switch 3 |
| --- | --- | --- | --- |
| ON | OFF | OFF | OFF |

#### Measurement results

Figure 18 - Linearity - Cf=1.6pF

| Buffer Structure | Widlar | Common collector |
| --- | --- | --- |
| Qin max | 3.4 pC | 2.5 pC |
| Vout max | 1.26 V (over 50 ) | 0.9 V (over 50 ) |
| Gain | 370 mV/pC | 358 mV/pC |
| Feedback cap | 1.35 pF | 1.4 pF |
| Non-linearity | 3.9‰ | 1.7‰ |

Table 10 - Characterization Cf=1.6pF

### Linearity measurement – Cf=0.8pF

#### Switch configuration

| Switch 0 | Switch 1 | Switch 2 | Switch 3 |
| --- | --- | --- | --- |
| OFF | ON | OFF | OFF |

#### Measurement results

Figure 19 - Linearity - Cf=0.8pF

| Buffer Structure | Widlar | Common collector |
| --- | --- | --- |
| Qin max | 1.73 pC | 1.4 pC |
| Vout max | 1.22 V (over 50 ) | 0.95 V (over 50 ) |
| Gain | 702 mV/pC | 681 mV/pC |
| Feedback cap | 0.71 pF | 0.73 pF |
| Non-linearity | 4.1‰ | 3.2‰ |

Table 11 - Characterization Cf=0.8pF

### Linearity measurement – Cf=0.4pF

#### Switch configuration

| Switch 0 | Switch 1 | Switch 2 | Switch 3 |
| --- | --- | --- | --- |
| OFF | OFF | ON | OFF |

#### Measurement results

Figure 20 - Linearity - Cf=0.4pF

| Buffer Structure | Widlar | Common collector |
| --- | --- | --- |
| Qin max | 0.87 pC | 0.72 pC |
| Vout max | 1.17 V (over 50 ) | 0.94 V (over 50 ) |
| Gain | 1350 mV/pC | 1310 mV/pC |
| Feedback cap | 0.37 pF | 0.38 pF |
| Non-linearity | 3.4‰ | 3.3‰ |

Table 12 - Characterization Cf=0.4pF

### Linearity measurement – Cf=0.2pF

#### Switch configuration

| Switch 0 | Switch 1 | Switch 2 | Switch 3 |
| --- | --- | --- | --- |
| OFF | OFF | OFF | ON |

#### Measurement results

Figure 21 - Linearity - Cf=0.2pF

| Buffer Structure | Widlar | Common collector |
| --- | --- | --- |
| Qin max | 0.47 pC | 0.41 pC |
| Vout max | 1.15 V (over 50 ) | 0.98 V (over 50 ) |
| Gain | 2460 mV/pC | 2390 mV/pC |
| Feedback cap | 0.203 pF | 0.21 pF |
| Non-linearity | 4.3‰ | 5.1‰ |

Table 13 - Characterization Cf=0.2pF

## Pedestal dispersion

Pedestal measurement have been made on both track-and-hold buffer structures (Common collector and Widlar, Figure 7). As a differential structure, we expected Widlar buffer to be significantly better than Common collector. Measurement of unit gain is shown on Figure 22.Measurement of Gain 10 is shown on Figure 23. Meassurement of pedestal dispersion is in Table 14.

| DC (V) | DC (V) |
| --- | --- |

Figure 22 - Pedestal measurement on Gain 1 output

| DC (V) | DC (V) |
| --- | --- |

Figure 23 - Pedestal measurement on gain 10 output

|  | Common collector | Widlar structure |
| --- | --- | --- |
| Gain 1 | 7,8 mV. | 0,9 mV. |
| Gain 10 | 16,9 mV | 1,1 mV. |

Table 14 - Pedestal measurement results

# ANNEX 1 – FLC_PHY2 pinout, bias and current

# ANNEX 2 – FLC_PHY2 versus FLC_PHY1 pinout

# ANNEX 3 - Preamp schematic

# ANNEX 4 – shaper gain 1 schematic

# ANNEX 5 – Shaper gain 10 schematic

# ANNEX 6 – Shapers OPAMP schematic

# ANNEX 7 – Widlar schematic