# FLC Front-end board Test

# Interface

## DC Supply

DC supply is provided through a 6-contact supply connector referenced as follow :

| Denomination | Reference |
| --- | --- |
| right angle pin header | Hirose DF1E – 6P – 2.5DS |
| Socket | Hirose DF1E – 6S – 2.5C |
| Socket Crimping contact | Hirose DF1E – 2022SC (AWG #22) |

The pinout of that connector is the following :

| Pin # | Description |
| --- | --- |
| 1 | High voltage diode bias (-200V) |
| 2 | VSS (-5V) |
| 3 | GND |
| 4 | VP6 (+6V) |
| 5 | GND |
| 6 | VM6 (-6V) |

High voltage diode bias should not flow any current but the diodes leakage current. Vss supplies all the front end and calibration switches. VP6 and VM6 supply all the line drivers and buffers.

## Read-out connection

Read-out connection is ensured by a AMP 68 pin-low-profile connector referenced 787254-1.

Pinout of the connection is given in the table below :

| Pin # | In/Out | Description |
| --- | --- | --- |
| 1, 35 | Out | Output 1, analogue differential (-,+ resp.) |
| 4, 38 | Out | Output 2, analogue differential (-,+ resp.) |
| 14, 48 | Out | Output 3, analogue differential (-,+ resp.) |
| 20, 54 | Out | Output 4, analogue differential (-,+ resp.) |
| 30, 64 | Out | Output 5, analogue differential (-,+ resp.) |
| 33, 67 | Out | Output 6, analogue differential (-,+ resp.) |
| 2, 36 | Out | Output 7, analogue differential (-,+ resp.) |
| 5, 39 | Out | Output 8, analogue differential (-,+ resp.) |
| 15, 49 | Out | Output 9, analogue differential (-,+ resp.) |
| 21, 55 | Out | Output 10, analogue differential (-,+ resp.) |
| 31, 65 | Out | Output 11, analogue differential (-,+ resp.) |
| 34, 68 | Out | Output 12, analogue differential (-,+ resp.) |
| 6, 40 | In | Hold, LVDS (-,+ resp.) – 1 : Track, 0:Hold |
| 8,42 | In | SRIN, LVDS (-,+ resp.), Multiplexer Shift register input |
| 9,43 | In | Reset, LVDS (-,+ resp.), Reset all shift register D-latch |
| 13, 47 | In | Clock, LVDS (-,+ resp.), Shift register clock |
| 23, 57 | In | Address, LVDS (-,+ resp.), lit a diode when settled to 1 (debug use) |
| 7, 41 | In | Vcalib, Analogue differential input |
| 25, 59 | In | Tcalib 1, LVDS (-,+ resp.) |
| 26, 60 | In | Tcalib 2, LVDS (-,+ resp.) |
| 10, 44 | In | Enable 1, LVDS (-,+ resp.), Enable cal channel 1&7 |
| 11, 45 | In | Enable 2, LVDS (-,+ resp.), Enable cal channel 2&8 |
| 12, 46 | In | Enable 3, LVDS (-,+ resp.), Enable cal channel 3&9 |
| 16, 50 | In | Enable 4, LVDS (-,+ resp.), Enable cal channel 4&10 |
| 17, 51 | In | Enable 5, LVDS (-,+ resp.), Enable cal channel 5&11 |
| 18, 52 | In | Enable 6, LVDS (-,+ resp.), Enable cal channel 6&12 |
| 27, 61 | Out | TYPE 0, LVDS, |
| 28, 62 | Out | TYPE 1, LVDS |
| 29, 63 | Out | TYPE 2, LVDS |
| 32, 66 | Out | TYPE 4, LVDS |
| 24, 58 | In | Gain 1 / Gain 10 selection switch |
| 22, 56 | Out | Multiplexers SR output |

NC pins are : 3, 19, 37, 53

Analogue differential outputs are driven by EL2140 and are matched to 120 Ohm. Common mode can be tuned independently of VFE chip output quiescent voltage.

Input LVDS are converted into CMOS with  CM26LS32 converters. LVDS output are converted from CMOS with 26LS31.

Diodes positions (Active pads highlighted)

| D16 | D13 | D10 | D7 | D4 | D1 | D124 | D121 | D118 | D115 | D112 | D109 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| D17 | D14 | D11 | D8 | D5 | D2 | D125 | D122 | D119 | D116 | D113 | D110 |
| D18 | D15 | D12 | D9 | D6 | D3 | D126 | D123 | D120 | D117 | D114 | D111 |
| D34 | D31 | D28 | D25 | D22 | D19 | D142 | D139 | D136 | D133 | D130 | D127 |
| D35 | D32 | D29 | D26 | D23 | D20 | D143 | D140 | D137 | D134 | D131 | D128 |
| D36 | D33 | D30 | D27 | D24 | D21 | D144 | D141 | D138 | D135 | D132 | D129 |
| D52 | D49 | D46 | D43 | D40 | D37 | D160 | D157 | D154 | D151 | D148 | D145 |
| D53 | D50 | D47 | D44 | D41 | D38 | D161 | D158 | D155 | D152 | D149 | D146 |
| D54 | D51 | D48 | D45 | D42 | D39 | D162 | D159 | D156 | D153 | D150 | D147 |
| D70 | D67 | D64 | D61 | D58 | D55 | D178 | D175 | D172 | D169 | D166 | D163 |
| D71 | D68 | D65 | D62 | D59 | D56 | D179 | D176 | D173 | D170 | D167 | D164 |
| D72 | D69 | D66 | D63 | D60 | D57 | D180 | D177 | D174 | D171 | D168 | D165 |
| D88 | D85 | D82 | D79 | D76 | D73 | D196 | D193 | D190 | D187 | D184 | D181 |
| D89 | D86 | D83 | D80 | D77 | D74 | D197 | D194 | D191 | D188 | D185 | D182 |
| D90 | D87 | D84 | D81 | D78 | D75 | D198 | D195 | D192 | D189 | D186 | D183 |
| D106 | D103 | D100 | D97 | D94 | D91 | D214 | D211 | D208 | D205 | D202 | D199 |
| D107 | D104 | D101 | D98 | D95 | D92 | D215 | D212 | D209 | D206 | D203 | D200 |
| D108 | D105 | D102 | D99 | D96 | D93 | D216 | D213 | D210 | D207 | D204 | D201 |

Diodes D1 to D108 are in the VFE0 part of the board, Diodes D109 to D216 are in the VFE1 part. On that first assembled version, only 4 front-end chip are mounted.

| VFE0 | Channel 1 | diodes 1 to 18 |
| --- | --- | --- |
|  | Channel 2 | diodes 19 to 36 |
|  | Channel 3 | diodes 37 to 54 |
|  | Channel 4 | diodes 55 to 72 |
|  | Channel 5 | diodes 73 to 90 |
|  | Channel 6 | diodes 91 to 108 |
| VFE1 | Channel 7 | diodes 109 to 126 |
|  | Channel 8 | diodes 127 to 144 |
|  | Channel 9 | diodes 145 to 162 |
|  | Channel 10 | diodes163 to 180 |
|  | Channel 11 | diodes 181 to 198 |
|  | Channel 12 | diodes 199 to 216 |

## Calibration circuit

There is only one Calibration chip mounted, that allows to calibrate read-out channel 1 and channel 2)

| VFE0 (left part) - MOUNTED |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Channel 1 | Channel 2 | Channel 3 | Channel 4 | Channel 5 | Channel 6 |
| D1 | D4 | D7 | D2 | D5 | D8 |
| D3 | D6 | D9 | D11 | D14 | D17 |
| D10 | D13 | D16 | D19 | D22 | D25 |
| D12 | D15 | D18 | D21 | D24 | D27 |
| D20 | D23 | D26 | D28 | D31 | D34 |
| D29 | D32 | D35 | D30 | D33 | D36 |
| D37 | D40 | D43 | D38 | D41 | D44 |
| D39 | D42 | D45 | D47 | D50 | D53 |
| D46 | D49 | D52 | D55 | D58 | D61 |
| D48 | D51 | D54 | D57 | D60 | D63 |
| D56 | D59 | D62 | D64 | D67 | D70 |
| D65 | D68 | D71 | D66 | D69 | D72 |
| D73 | D76 | D79 | D74 | D77 | D80 |
| D75 | D78 | D81 | D83 | D86 | D89 |
| D82 | D85 | D88 | D91 | D94 | D97 |
| D84 | D87 | D90 | D893 | D96 | D99 |
| D92 | D95 | D98 | D100 | D103 | D106 |
| D101 | D104 | D107 | D102 | D105 | D108 |

| VFE1(Right part) – NOT MOUNTED |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Channel 7 | Channel 8 | Channel 9 | Channel 10 | Channel 11 | Channel 12 |
| D109 | D112 | D115 | D110 | D113 | D116 |
| D111 | D114 | D117 | D119 | D122 | D125 |
| D118 | D121 | D124 | D127 | D130 | D133 |
| D120 | D123 | D126 | D129 | D132 | D135 |
| D128 | D131 | D134 | D136 | D139 | D142 |
| D137 | D140 | D143 | D138 | D141 | D144 |
| D145 | D148 | D151 | D146 | D149 | D152 |
| D147 | D150 | D153 | D155 | D158 | D161 |
| D154 | D157 | D160 | D163 | D166 | D169 |
| D156 | D159 | D162 | D165 | D168 | D171 |
| D164 | D167 | D170 | D172 | D175 | D178 |
| D173 | D176 | D179 | D174 | D177 | D180 |
| D181 | D184 | D187 | D182 | D185 | D188 |
| D183 | D186 | D189 | D191 | D194 | D197 |
| D190 | D193 | D196 | D199 | D202 | D205 |
| D192 | D195 | D198 | D201 | D204 | D207 |
| D200 | D203 | D206 | D208 | D211 | D214 |
| D209 | D212 | D215 | D210 | D213 | D216 |

OUT Ch.1

OUT Ch.2

OUT Ch.3

OUT Ch.4

OUT Ch.5

OUT Ch.6

Vcalib

Tcalib2

Enable Ch. 1

Enable Ch. 2

Enable Ch. 3

Enable Ch. 4

Enable Ch. 5

Enable Ch. 6

Calibration circuit

1

OUT Ch.7

OUT Ch.8

OUT Ch.9

OUT Ch.10

OUT Ch.11

OUT Ch.12

Calibration circuit

2

Tcalib1

T

T

V

V